SETUP OF MICROLUMI XSi WITH INJECTOR(S)

CAUTION:

- TURN OFF THE MICROLUMI-XS BEFORE INSTALLING OR REMOVING ANY CABLES (POWER CABLE, INJECTOR CABLES).

- ALWAYS USE THE ORIGINAL INJECTOR CABLES, USE OF OTHER CABLES WILL DAMAGE THE UNITS AND WILL VOID WARRANTY.

- ALWAYS EVACUATE THE INJECTORS IF THEY ARE NOT GOING TO BE USED FOR PROLONGED PERIOD.

- FOR REPLACEMENT CABLES OR ANY OTHER PARTS, CONTACT HARTA INSTRUMENTS.

CONNECTING THE INJECTORS TO THE LUMINOMETER:

- Turn OFF the luminometer.

- Black tubing with two fittings :

  - Connect Blue ¼-28 fitting to “OUT” port of the valve (see fig 1, “A”)

  - Connect Male Luer fitting to the back of the MicroLumi XS.

- Clear tubing with one fitting:

  - Connect the ¼-28 fitting to “IN” port of the valve (see fig 1, “B”)

  - Dip the other end into the fluid to be injected.

- Attach cables from the back of machine to it’s corresponding injectors.

REPLACING THE SYRINGE:

Use only Becton Dickinson Part# 309628 1 ml syringe, use of any other type of syringe may not be accurate and may cause damage to the machine. Additional syringes available from Harta Instruments, Inc.

- From Lumiterm, evacuate the injector(s)

- At the end of evacuate cycle, the injector arm will be down.

- Release the plunger holder lever

- Unscrew the syringe

- Screw-in a new syringe

- Make sure the syringe tube wings are paralel to the front panel, to prevent injector arm to interfere with the syringe tube (see fig. 2, “A”)

- Tighten the plunger holder lever (see fig. 2, “B”)

REPLACING TUBING:

- From Lumiterm, evacuate the injector(s)

- Turn OFF the luminometer

- Remove the fitting from the valve body

- Remove the Luer from the back of MicroLumi XS

- Replace tubing

- Additional Fittings are available from Harta Instruments Inc

PRIMING THE INJECTOR:

CAUTION:

FOR RELIABLE OPERATION, IT IS IMPORTANT TO ENSURE THERE’S NO AIR COLUMN IN THE PIPELINE. FOLLOW THE FOLLOWING PROCESS:

- Place an empty microplate in the MicroLumi XS tray (use microplate with microwell volume > 200 uL), close the tray

- Remove luer(s) from the back of luminometer, and loop it back to it’s supply container

- From “Options”, select columns A, B and C

- Select “Flash”

- for each active injectors:

  - Set “Volume” to 100 uL

  - “Delay” = 0

  - “Read” = 1

  - Shuttr speed = 100

  - Kinetic delay = 0

- Click “OK”

- Click “Start”

- If there’s air column in the syringe or the tubings, repeat the process

- When at the end of the cycle, the syringe is free of air column, re-install the luer to the back of luminometer

- From “Options” select column A

- Click “OK” and “Start”

FLUSHING THE INJECTOR:

- From Lumiterm, evacuate the injector(s)

- Place an empty microplate in the MicroLumi XS tray (use microplate with microwell volume > 200 uL), close the tray

- Dip the input tubing into rinsing / flushing solution

- Proceed to “Read” the whole plate in Flash mode with the same parameters as in step 3 of PRIMING

- Repeat as necessary

- Evacuate

USER REPLACEABLE PARTS LIST:

- 1 ml syringe					HDSI-101

- Output tubing assy				HDSI-102

- Input tubing assy				HDSI-103

- ¼ - 28 barbed fitting				HDSI-104

- Male luer barbed fitting				HDSI-105

- Control cable					HDSI-201

Fig. 1						  Fig. 2

A

B

A

B

HARTA INSTRUMENTS, INC.

8 RUSSELL AVE, UNIT 106, GAITHERSBURG, MD 20878

PHONE:301-948-8868.       EMAIL: HARTA@EROLS.COM

WWW.HARTACORPORATION.COM